## Slide 1: The current status of “masha” setup
Krupa L.
Flerov Laboratiory of Nuclear Reactions, JINR, Dubna, Russia

## Slide 2: Elements obtained in Dubna first
18O + 243Am→261Lr* →256Lr + 5n
249Cf+ 15N → 264Db* → 260Db + 4n
249Cf + 18O→263Sg + 4n
Cn… Nh, Fl, Mc, Ts, Og.
To be continued…

## Slide 3: FLNR basic installations
U-400
U-400M
DC-280
IC-100

## Slide 4
Special view

## Slide 5
Stable elements
(Continent)
Transuranium el. (Peninsula)
Super Heavy el.
(Island)
Cold fusion reactions
Hot fusion reactions
The stability map

## Slide 6
Mass Analyzer of Super Heavy Atoms
MASHA
“MASHA” at the beam line of the cyclotron U-400M, FLNR, JINR

## Slide 7
The present status of the mass-spectrometer MASHA (Mass-Separator of Supper Heavy Atoms) designed for determination of the masses of superheavy elements.
Results of the test experiments on the mass-spectrometer MASHA aimed at the study of evaporation residua in the reactions 40Ar+144Sm and 40Ar+166Er.
Production and spectroscopic investigation of neutron-rich isotopes, produced in mutinucleon transfer reactions, near the neutron N = 126 and N=152 shell closures. Test experiments.
Determination of masses of super heavy elements in the experiments on synthesis of 112 and 114 elements by using the reactions 48Ca+238U and 48Ca+242Pu. The beam energy is Ebeam ~ 5 MeV/n.
Contents

## Slide 8
The proposed setup is a combination of the so-called ISOL method of synthesis and separation of radioactive nuclei with the classical method of mass analysis, allowing mass identification of the synthesized nuclides in the wide mass range (A = 1-450 u).
General ion-optical parameters:
Range of energy variation, keV 15-40
Range of Br variation, Tm 0.08-0.5
Mass acceptance, % +/-2.8
Angular acceptance, mrad +/-14
Diameter the ion source exit hole, mm 5.0
Horizontal magnification at F1/F2 0.39/0.68
Mass dispersion at F1/F2, mm/% 1.5/39.0
Linear mass resolution at F1 75
Mass resolution at F2 1300
1 – Target box with hot catcher;
2 – Ion source;
3 – Mass separator;
4 – DAQ in the focal plane.
Mass
Analyzer of
Super
Heavy
Atoms

## Slide 10
Material of the catcher – flexible graphite, density of 1 g/cm3,
a thickness of 0.6 mm and shaped as a 30 mm diameter disk.
Operating temperature of hot catcher – 1800-2000оС.
Delivery time of nuclides to the ECR ion source < 1.8 s.
ECR
ion source
Hot catcher
Target
Beam line
Recoil transport
Hot catcher scheme

## Slide 11
Ion source + Hot catcher (top view)
Nuclear reaction products escape from the target, pass through the separating foil, and are stopped in the graphite absorber. In the form of atoms, the products diffuse from the graphite absorber to the vacuum volume of the hot catcher and, moving over the pipeline, reach the ECR source, where are ionized to charge state Q = +1 and accelerated with the aid of the three electrode system.

## Slide 12
242Pu rotating target before and after irradiation

## Slide 13: Hot Catcher
Heating up to 1800 oC
In assembly on a flange
Under microscope
Sheets
Porosity coefficient – 75%

## Slide 14
Focal plane silicon multi strip detector
Configuration – well type
Number of the focal strips – 192 (step – 1.25 mm)
Number of the back side strips – 160 (step – 5 mm)
Total efficiency – more than 90%
TIMEPIX detector
An array of 256x256 square pixels of pitch size 55mm for full sensitive area 14x14 mm2. Silicon sensor of 300 mm width.

## Slide 15: Data acquisition system
22 charge-sensitive preamplifier, 16 channels.
22 multiplexers-amplifiers designed at FLNR JINR.
5 digitizers from XIA (PXI standard), 16 channels, 250 MHz, 12-bits.
Energy resolution ~ 25 keV.
After amplification and multiplexing, three outputs from the multiplexer are used: alpha, fragment, and digital channels. The digital channel includes the information on the source of the input signal. These outputs are connected to the XIA multichannel high-speed digitizers.

## Slide 16
PXI chassis with XIA digitizers, NI controller and crate with multiplexors (left photo). Examples of signals from one strip (right pictures).

## Slide 17: Energy measurement
Directly:
By help of 2 pick-up detectors with 2 pre-amplifiers designed at FLNR, JINR.
Using scatterring beam on gold foil (45o):
With the aid of 2 MCP detectors, distance - 1210 mm work as a Time-of –Flight and 1 latter Silicon detector work as TKE.
TOF method
Electronics: (PXI standard)
2 High-speed digitizers from Agilent, 2 channels, 2 GHz, 10-bits.
2 TDC from Agilent, 5 ps resolution, 12 channels.

## Slide 18
First test experiments
40Ar+144Sm -> Hg, Ebeam = 5-7 MeV/n
40Ar+166Er -> Rn, Ebeam = 5-7 MeV/n
40Ar5+, current ~ 100 pnA

## Slide 21
40Ar+144Sm→184-xnHg

## Slide 22
40Ar+166Er→166-xnRn

## Slide 23: Theoretical calculations, coupling channel model

## Slide 24
Multi nucleon transfer reactions
Calculation for 48Ca+232Th
at the energy 370 MeV
The composite target 232Th was used in the experiments. An alcohol solution of thorium nitrate was deposited into flexible graphite (thickness = 0.6 mm) and annealed at the temperature 1900°C. Finally target as ThO2 was made and bombarded by beam of 48Ca with energy Ebeam = 7.3 MeV/nucl.
Composed target
232Th + graphite

## Slide 25: Test experiments on production of Rn and Hg isotopes in the multi nucleon transfer reactions by using the TIMEPIX detector
This camera operates as an active nuclear emulsion enabling the detection of heavy charged particles, such as fission fragments, and α-particles, β-particles, X and γ -rays in a wide linear dynamic range.

## Slide 26: Timepix: HCP’s: a’s, fission fragments

## Slide 28
Radon isotopes
229Rn
231Rn
232Rn
230Rn

## Slide 29: Thank you for attention